Technical specification of Call for proposals No SZTE/2021/PSZ11834 for Design, develop, and implement a high pressure gas target system (HPGTS) for deuteron acceleration

University of Szeged is launching a call for proposal for design, develop and implement a high pressure gas target system. University of Szeged is a distinguished user of ELI-ALPS, Szeged, to study and develop a laser based neutron source. The R&D project aims at the development a gas target system, which is capable for accelerating deuterium ions with the use of few cycle laser pulses. The gas target system will be installed and run at ELI-ALPS, Szeged, Hungary, as part of the user (laser neutron source) beamline run by University of Szeged.

I. Research task

Design, develop, implement, and install a HPGTS which shall be operational with the SEA laser parameters (10 Hz, <12 fs pulse duration, 30mJ). As a best effort, it should aim at the operation at 1kHz repetition rate.

II. Technical specifications and scoring

For the specification of the HPGTS unit:

Scored parameters:

| # | Specification | Weight |
| --- | --- | --- |
| II.1 | Cut-off energy of deuteriums: minimum: 400keV, MfV: 1MeV | 10 |
| II.2 | Total number of deuteriums above 50keV per shot: minimum: 10^7; MfV: 10^10 | 15 |
| II.3 | Repetition rate: minimum: 10Hz, Mfv: 100Hz | 10 |
| II.4 | Best effort – repetition rate minimum 100Hz Mfv: 1kHz | 5 |
| II.5 | Duration of continuous operation without human interaction Minimum: 1 h;MfV: 4 hours | 5 |
| II.6 | Duration of the project (starting from contracting): Maximum: 15 months, Mfv: 10 months | 5 |
| II.7 | Net total cost | 50 |

MfV: Most favourable value. Better specifications than MfV would not be scored higher than that of the Mfv value.

Best effort values: upon implementation reasonable efforts shall be devoted and reported to achieve the offered best effort values. In the lack of such efforts and documentation, up to 5%-5% of the total R&D cost would be withheld. The best effort value (II.4) shall exceed the offered normal value (II.3) by at least 15 %.

Calculation method:

- For II.1-II.6: in direct proportion to the most favourable value, corrected by extreme value. For parameters that reach or exceed the best value, maximum points shall be given.

Equations:

II.1, II.2,II.5

Pscored=(VALUEOffered-VALUEMax)/(VALUEMfv-VALUEMax)*WEIGHT

II.3, II.4

Pscored=(VALUEOffered-VALUEMin)/(VALUEMfv-VALUEMin)*WEIGHT

- For II.7: in inverse proportion to the lowest offer as best value

Equation:

Pscored=VALUEOfferedBest/VALUEOffered*WEIGHT

Total score: the sum of the particular scores, rounded to 2 digits.

Boundary conditions and requirements

II. 8 Residual pressure in the interaction chamber: not higher than 5×10^-5 mbar.

II. 9 The system shall include:

- high pressure gas nozzles (plus at least 2 spare ones)

- necessary gas handling system (compressor, booster, etc.);

- monitoring and control units and interfaces (excluding computer)

III. Project proposal

Please, prepare one project proposal for a HPGTS meeting the specifications and requirements listed in section II. The proposal shall be reasonably succinct, but shall not be longer than 30 pages. The proposal and tables, figures, and equations, etc. contained in it shall be legible (at least Times Roman 10 or similar).  It shall contain the state of the art of the given field, a brief description of the scientific problem or technical bottleneck, and a short description of the planned research activity, and literature / references, as well as the scientific and/or engineering experience and/or activity of the bidder. The technical proposal shall be written in sufficient detail so that the feasibility of the technical specifications detailed in chapter II. can be established by an expert body. The proposal is expected to provide a feasible research strategy plan. More detailed solutions and know-hows are expected to be detailed by the winner in the frame of the TDR (DL2).

The Requesting Party has defined some key deliverables and documents which must be definitely included in the research plan. The proposal should include the technical specifications and/or operational parameters of the deliverables (DL). It shall contain the scheduling and specifications of preparation of the components as well as the preliminary list of accessories for control-command and diagnostics. In addition, please define further milestones and/or deliverables, which could be served as a half-a year year monitoring points of the project. Please define the suggested date of delivery or milestone for each cases.

DL.1: Preliminary studies (if any; please specify); (max 5%).

DL.2: Technical Design Report (template is provided); (10%).

DL.3: Delivery to Szeged (20%)

DL.4: Installation (max 65%) up to the scored values (II.1-3 and II.5)

DL.5: Demonstration of the best effort values (II.4) (5%)

DL.6: Training and full documentation (10%)

IV. Execution of the R&D project

The R&D work of the task shall be executed on the premises of the contracted party at no additional cost. The R&D work shall be regarded as a joint research activity between the contracted party and University of Szeged, including any form of scientific communication (presentations, papers, etc.). The contracted party shall notify University of Szeged about the start and foreseeable duration of major experimental campaigns related to the R&D work contract and shall be committed to providing access, supervision and training for University of Szeged scientific personnel on the premises during these campaigns.

V. Bidder’s proposal shall contain the followings

- Please refer to Ref No SZTE/2021/PSZ11834

- Customer name: University of Szeged, H-6720 Dugonics tér 13., Szeged, Hungary

- Name and company data of the Bidder

- The content of the proposal according to section II.

- Please declare that you have sufficient human capacity in case of performing multiple tasks as well.

- Declaration of the professional experience in laser-driven particle acceleration. Content of the declaration – at least one of the followings:

- Name and degree of at least one professional hired by the Bidder within the last 12 months;

- At least 5 item list of scientific dissemination with the Bidders affiliation from the last 36 months.

- The cost of the R&D work shall incorporate all the related expenses, with special attention to the followings:

  - the total expenses of the R&D work shall cover all the costs for the necessary theoretical and experimental works, of all intermediate pilot set-ups, of the final version, of the device control and command system as well as person power, labor spaces, service expenses, etc.

  - shipment and on-site installation;

  - the rights of use (including IP rights – see section VI. below)

  - spare parts for the most critical ones of the system. Please provide a list of them.

  - the cost of one year of on-site service and warranty.

  - Please do not list the bank charge (to be paid by customer).

- Please provide the necessary diagnostics for the acceptance tests (may not be part of the delivered items, unless otherwise specified). Please note that the University may provide measurement devices, too.

- Terms of delivery (according to section III, see the’suggested date of delivery or milestone’ subsection)

- Terms of payment. Our University prefers to pay with wire transfer.

- Quote validity and/or quote issue date. The quotation should be preferably valid for 60 days, due to excessive evaluation period, but at least for 30 days. Quote validity day should be not earlier than the deadline of the procedure

- Please make sure that the proposal is signed by the authorized representative of your establishment.

- Also, please fill out, sign and submit the attached Summary Sheet

VI. Intellectual Property

The details relating to the IP shall be specified in the R&D contract between the Requesting Party and the bidder.

VII. Parameters of the SEAL laser

| Energy >40mJ, 0.9% rms over 19 hours | Spectrum |
| --- | --- |

Note, that the at least 15% of the laser energy are lost from the laser to the entrance of the HPGTS unit due to various diagnostics and beam transport. The exact value are to be determined at the time of contracting.

Pulse duration

Pointing stability:

|  | Ponting x direction (RMS): 2.5 µrad Ponting y direction(RMS): 1.2 µrad |
| --- | --- |

Near field

Date: 05th of October, 2021

Please upload your offer and the Summary Sheet on the website of the Call for the Proposal, No SZTE/2021/PSZ11834