Technical specification of Call for proposals SZTE/2022/PSZ12789

for Dosimetry Services for Extremity Doses

Service name: Dosimetry services for 25 ring dosimeters that measures X-rays, gamma rays and beta radiation and 3 sterile containers

Required service:

- service period: 12 months,

- dose data evaluation bi monthly

- 25 dosimeters

Specifications:

Requirements for service provider:

- The service provider is requested to be an accredited laboratory. The accreditation should be in accordance with the ISO/IEC 17025 international standard.

- The service provider is requested to participate in national and international inter-comparison studies.

Form of certificate: Tenderer must attach a simple statement about the participation in the studies.

Content of the declaration: please provide a list of the dates of the studies mentioned above.

Technical specifications:

- Passive ring dosimeters that measures X-rays, gamma rays and beta radiation.

- ThermoLuminescent Dosimeter (TLD) made from lithium fluoride (LiF).

- The ring fits the finger, adapting to its shape, is light and easy to wear, and thin enough not to hinder the movements.

- Can be worn under gloves without risk of tearing.

- The TLD is sealed inside the body of the ring to ensure it remains in place as long as the ring is on the finger.

- Operating and storage temperature not to have effect on detector sensitivity up to 40 °C, as well as exposure to light.

- Easy to clean and disinfect.

- The disinfection methods to be validated in the hospital environment

- Personal dose equivalent: Hp(0.07)

- Dose range: from 0.1 mSv to 10 Sv

- Minimum reporting value: 0.1 mSv

- Linearity response: from 0.10 mSv to 10 Sv Standard deviation < 10%

- Energy and angular response: ± 60° from 15 keV to 1.2 MeV

Shipping and handling cost of the dosimeters have to be included in the offer.

Please indicate shipping/handling costs in the offer.

Formal requirements of the offers. Please include these data in your proposal:

▪ Please refer to SZTE/2022/ PSZ12789

▪ Name of the bidder company, address. Please note that the bidder company, the contracting company and the invoicing company must be the same. If you have more than one company/ subsidiaries please provide the name of the invoicing company. (In case the bidder company outsources the invoicing, please notify us before the tender deadline)
▪ The NET price for the full service + all the expenses (e.g. shipping to Szeged, Hungary; handling etc.) in one offer.

- Please note that quoted net price (subtotals and totals) must match the prices listed in the final invoice to be issued upon delivery.

  - Please note that in case shipping/handling cost is not clearly listed in the offer, we will assume that the total cost includes shipping/handling cost, and no further shipping/handling cost can be charged.

  - (In case a bidder needs to pay additional taxes, fees, licenses etc after the products/services offered, they are kindly requested to be aware of that and consider these extra cost when calculating the offer net total price.)

  - Please do not add bank charge as this will be paid by customer.

  - Please list the net total sum

- Terms of delivery

- Terms of payment (such as: net 30 or net 60 or prepayment). Our University prefers to pay with wire transfer.

- Quote validity and/or quote issue date.

▪ The quotation should be preferably valid for 60 days, due to excessive evaluation period, but at least for 30 days.
▪ Quote validity day should be not earlier than the deadline of the procedure.
▪ Quote issue date should not be earlier than the starting date of the call for proposals.

- Please include the specification in your offer (so that it can be clearly evaluated whether it matches all the specified features).
Warranty and guarantee terms if applicable.

Thank you very much in advance for your cooperation!

Version date: 19/December/2022.

Please upload your offer via the website of the Call for the Proposal, 
No. SZTE/2022/PSZ12789