PB13357 specification — gas analyzer

•	CO2 concentration measurements between 0-100 V/V%

•	CO concentration measurements between 0-100 V/V%

•	H2 concentration measurements between 0-50 V/V%, but no permanent instrument damage should be caused by 100 V/V% H2.

•	CH4 concentration measurements between 0-50 V/V%

•	C2H4 concentration measurements between 0-20 V/V%

•	The concentration of a given gas should be independent on the concentration of the other components that are measured with other sensors — in case of cross-sensitivity, the instrument should compensate to it automatically.

•	The minimal gas flow rate is 1l/minute or lower. The instrument should sample the analyzed gas portion from a pipe in which the gas flow rate is between 20-200 l/min

•	The measured values should be independent on the gas flow rate (above the minimum limit).

•	PC communication possibility (RS-485 is preferred, but equivalent connection types can be accepted) — data transfer possibility to a NI CompactRIO system.

•	A minimum 10m long cable is supplied for the communication with the PC.

•	The instrument calibration should not be performed more frequently than 1 month, with maximum 2 gas mixtures and a zero-point. The calibration must be performable by the users.

•	If necessary, temperature correction is performed automatically either by the instrument, or in the software.

•	The error of the measured gas concentrations should be below 0.1 %

•	No regular maintenance requirement.

•	The instrument should be resistant to trace amount of water vapor (dew point maximum 4 °C), and it should not cause more than 1rel% uncertainty in the measured values.

•	Hungarian/English user manual (including notes for calibration, maintenance, error messages etc.)

•	Warranty: minimum 2 years

•	Delivery time: 2 months